## Slim Edge Measurements
	Alic Spellman	02/27/19
	New Sensors			“W0X-XX-d” means that the sensors were dropped in the gelpack
	Slim Edge Measurements			“Defects” refers to any maximally protruding defect in the region of measurement.
				“Back Focus” refers to changing the focal plan to capture the edge of any background wafer protrusion, aka below/behind the wafer surface.
			(Micrometers)			SCIPP	CNM		bad channels
Rank	CNM table number	Sensor	Sector A	Center	Sector B	AVG Dis.	Distance	Vfd	Sector A	Sector B	Vbd CNM	Vbd	I(Vbd) [uA]	Comment
	11	W07-S2	301.95	427.82	491.31	396.63	431	23	143, 173, 16, 150, 198	1, 21, 71, 103, 217, 102	60	80	8.33	smudges
5	12	W07-S3	481.91	355.245	255.39	368.65	382	23	194	165, 199, 88, 200	380	130	3.85	scratches, GR debri
		S3-Defects	506.92
3	13	W07-S5	439.53	314.56	214.81	327.17	342	23	117, 127, 128	43, 156, 164, 212	420	170	5.25	GR scratch
		S5-Defects	456.64
6	14	W07-S6	459.64	331.86	229.5	344.57	358	23	192, 194	13	150	100	4.37
		S6-Defects
	15	W09-S2	210.49	406.07	508.65	359.57	436	30	159, 197, 8, 44, 188	73	240	180	1.51
		S2-Back Focus	290.29			399.47
4	16	W09-S5	431.26	321.98	253.52	342.39	345	30	27, 86, 100, 116	100	120	120	1.43
		S5-Defects	561.99
	3	W02-S1-d	234.12	481.26	598.8	416.46	509	20	4	24	50	40	8.93	GR chipped
		S1-Back Focus	305.16			451.98
	5	W02-S4-d	428.83	319.6	233.48	331.155	349	23	35, 221, 138	29, 77, 98	150	50	11.39	GR chipped
	7	W03-S6-d	279.8	279.09	376.87	328.335	304	18	44, 186, 200	149, 213, 36, 58	280	60	11.27	GR chipped
2	8	W04-S3-d	456.9	340.31	241.99	349.445	418	25	97, 187, 227, 52, 172	49, 123, 86	350	370	0.63
1	9	W04-S6-d	215.14	289.91	392.36	303.75	313	18	51, 218	40, 48, 90	480	420	2.49	edge pondbads have artifacts
		S6-Back Focus	287.82			340.09
	10	W07-S1-d	214.91	273.51	364.56	289.735	300	23	101, 207, 209, 247, 110, 192, 208	88	>500	50	3.9	GR chipped
		S1-Back Focus	228.58			296.57
20190227 Sensors
IV Curve
Bias (Voltage)
Leakage Current (A)

## IV Curve  Data
Tested: Alic Spellman (UCSC)
“-d” means that sensors shifted in gelpack during shipping. .Some damage was done.
W09-S5			W07-S6			W09-S2			W07-S2			W07-S3			W07-S5				W02-S1-d			W02-S4-d			W03-S6-d			W04-S3-d			W04-S6-d			W07-S1-d			W07-S6 (03/11)			W02-S1 (03/11)
V	I		V	I		V	I		V	I		V	I		V	I			V	I		V	I		V	I		V	I		V	I		V	I		V	I		V	I
0	5.00025E-08		0	2.93712E-07		0	3.5354E-08		0	3.51471E-07		0	2.54354E-07		0	3.18295E-07			0	1.86808E-07		0	2.09322E-07		0	1.81382E-07		0	4.00717E-08		0	4.17505E-08		0	3.26793E-07		0	2.99965E-07		0	9.52249E-08
5	1.80037E-07		5	1.5414E-06		5	1.27279E-07		5	4.63213E-06		5	1.16906E-06		5	1.47095E-06			5	1.58183E-06		5	3.63805E-06		5	2.72762E-06		5	1.07114E-07		5	1.28819E-07		5	1.4291E-06		5	1.45096E-06		5	6.48753E-07
10	2.27558E-07		10	2.78942E-06		10	1.82398E-07		10	6.73681E-06		10	2.06337E-06		10	2.48291E-06			10	3.66141E-06		10	6.60425E-06		10	5.86E-06		10	1.5748E-07		10	1.85694E-07		10	2.40609E-06		10	2.53632E-06		10	1.12245E-06
15	2.5817E-07		15	3.26643E-06		15	2.16792E-07		15	7.37049E-06		15	2.4363E-06		15	3.02981E-06			15	5.51929E-06		15	8.16467E-06		15	7.35395E-06		15	1.97362E-07		15	2.3598E-07		15	2.9026E-06		15	3.17761E-06		15	1.51272E-06
20	2.6753E-07		20	3.32684E-06		20	2.35896E-07		20	7.51392E-06		20	2.48603E-06		20	3.14599E-06			20	6.7074E-06		20	8.98159E-06		20	8.42784E-06		20	2.09306E-07		20	2.57381E-07		20	3.04682E-06		20	3.32652E-06		20	1.61574E-06
25	2.78059E-07		25	3.36337E-06		25	2.51798E-07		25	7.5355E-06		25	2.51424E-06		25	3.18164E-06			25	7.40972E-06		25	9.59423E-06		25	9.05725E-06		25	2.14267E-07		25	2.68714E-07		25	3.13993E-06		25	3.35906E-06		25	1.71563E-06
30	2.84665E-07		30	3.39289E-06		30	2.61821E-07		30	7.56604E-06		30	2.53853E-06		30	3.21417E-06			30	7.92357E-06		30	1.04294E-05		30	9.69274E-06		30	2.17777E-07		30	2.75594E-07		30	3.2441E-06		30	3.38501E-06		30	1.75392E-06
35	2.88588E-07		35	3.41739E-06		35	2.70133E-07		35	7.57673E-06		35	2.55964E-06		35	3.24185E-06			35	8.2785E-06		35	1.05409E-05		35	9.8908E-06		35	2.19978E-07		35	2.80805E-07		35	3.37054E-06		35	3.40755E-06		35	1.86425E-06
40	2.91237E-07		40	3.43924E-06		40	2.72192E-07		40	7.56698E-06		40	2.57863E-06		40	3.27061E-06			40	8.92909E-06		40	1.08259E-05		40	1.02714E-05		40	2.23112E-07		40	2.88962E-07		40	3.51198E-06		40	3.42835E-06		40	2.04678E-06
45	2.96088E-07		45	3.45931E-06		45	2.81384E-07		45	7.56191E-06		45	2.59719E-06		45	3.29725E-06			45	9.48041E-06		45	1.11328E-05		45	1.05354E-05		45	2.25354E-07		45	2.99643E-07		45	3.68804E-06		45	3.44716E-06		45	2.10983E-06
50	3.03208E-07		50	3.47996E-06		50	2.89714E-07		50	7.59859E-06		50	2.61573E-06		50	3.32221E-06			50	9.96752E-06		50	1.13855E-05		50	1.07493E-05		50	2.2772E-07		50	3.10984E-07		50	3.89562E-06		50	3.46473E-06		50	2.19303E-06
55	3.12977E-07		55	3.49856E-06		55	2.97272E-07		55	7.62486E-06		55	2.63352E-06		55	3.34647E-06			55	1.06827E-05		55	1.19169E-05		55	1.10267E-05		55	2.30828E-07		55	3.23964E-07		55	4.13259E-06		55	3.48018E-06		55	2.35553E-06
60	3.26003E-07		60	3.51852E-06		60	3.03426E-07		60	7.69001E-06		60	2.65246E-06		60	3.37334E-06			60	1.12709E-05		60	1.24641E-05		60	1.12706E-05		60	2.31641E-07		60	3.34229E-07		60	4.39443E-06		60	3.49645E-06		60	2.44028E-06
65	3.51259E-07		65	3.54241E-06		65	3.1315E-07		65	7.78739E-06		65	2.67534E-06		65	3.40008E-06			65	1.18989E-05		65	1.33097E-05		65	1.17136E-05		65	2.33619E-07		65	3.45789E-07		65	4.70648E-06		65	3.51351E-06		65	2.54432E-06
70	3.97164E-07		70	3.5719E-06		70	3.2018E-07		70	7.93244E-06		70	2.70167E-06		70	3.424E-06			70	1.26735E-05		70	1.39329E-05		70	1.20776E-05		70	2.35324E-07		70	3.56666E-07		70	5.05572E-06		70	3.52859E-06		70	2.69516E-06
75	4.68938E-07		75	3.60771E-06		75	3.3006E-07		75	8.0873E-06		75	2.73E-06		75	3.45091E-06			75	1.38672E-05		75	1.49443E-05		75	1.26683E-05		75	2.36869E-07		75	3.69713E-07		75	5.45136E-06		75	3.54531E-06		75	2.82995E-06
80	5.53069E-07		80	3.6536E-06		80	3.40194E-07		80	8.33396E-06		80	2.76113E-06		80	3.47727E-06			80	1.50238E-05		80	1.59681E-05		80	1.32182E-05		80	2.38119E-07		80	3.80916E-07		80	5.89967E-06		80	3.56333E-06		80	3.02563E-06
85	6.58641E-07		85	3.72403E-06		85	3.69914E-07		85	8.57515E-06		85	2.79524E-06		85	3.51659E-06			85	1.6032E-05		85	1.72939E-05		85	1.39337E-05		85	2.39852E-07		85	3.91462E-07		85	6.4259E-06		85	3.58497E-06		85	3.17709E-06
90	7.66622E-07		90	3.83334E-06		90	4.24399E-07		90	8.88588E-06		90	2.83644E-06		90	3.54387E-06			90	1.75114E-05		90	1.89009E-05		90	1.44771E-05		90	2.41308E-07		90	4.03125E-07		90	6.99115E-06		90	3.60769E-06		90	3.3558E-06
95	8.53997E-07		95	4.02711E-06		95	4.64391E-07		95	9.24457E-06		95	2.88731E-06		95	3.57202E-06			95	1.89575E-05		95	1.99835E-05		95	1.51557E-05		95	2.42562E-07		95	4.14256E-07		95	7.64022E-06		95	3.63298E-06		95	3.54721E-06
100	9.17035E-07		100	4.37319E-06		100	4.9519E-07		100	9.64098E-06		100	2.94755E-06		100	3.60108E-06			100	1.99151E-05		100	1.99831E-05		100	1.59975E-05		100	2.44438E-07		100	4.2412E-07		100	8.32466E-06		100	3.66108E-06		100	3.7078E-06
105	1.03055E-06		105	4.98666E-06		105	5.1143E-07		105	1.0134E-05		105	3.02411E-06		105	3.64327E-06			105	1.9983E-05		105	1.99842E-05		105	1.68517E-05		105	2.46934E-07		105	4.37689E-07		105	9.15335E-06		105	3.69412E-06		105	3.9312E-06
110	1.13204E-06		110	5.97706E-06		110	5.39108E-07		110	1.06456E-05		110	3.11491E-06		110	3.67392E-06			110	1.99851E-05		110	1.99827E-05		110	1.77347E-05		110	2.483E-07		110	4.52308E-07		110	1.27618E-05		110	3.73266E-06		110	4.14004E-06
115	1.25167E-06		115	7.65463E-06		115	5.61797E-07		115	1.12692E-05		115	3.21967E-06		115	3.71573E-06			115	1.99837E-05					115	1.8496E-05		115	2.50519E-07		115	4.65627E-07		115	1.99647E-05		115	3.78111E-06		115	4.37246E-06
120	1.43328E-06		120	1.0397E-05		120	6.43671E-07		120	1.1929E-05		120	3.32572E-06		120	3.75289E-06			120	1.99829E-05					120	1.95351E-05		120	2.5193E-07		120	4.77789E-07		120	1.9982E-05		120	3.84292E-06		120	4.55713E-06
125	1.63288E-06		125	1.50311E-05		125	7.35569E-07		125	1.26488E-05		125	3.45234E-06		125	3.79283E-06					125	1.99834E-05		125	2.54122E-07		125	4.90564E-07		125	1.99936E-05		125	3.91973E-06		125	4.78402E-06
130	1.83895E-06		130	1.99815E-05		130	6.30495E-07		130	1.99636E-05		130	3.61538E-06		130	3.85154E-06					130	1.99824E-05		130	2.56127E-07		130	5.03356E-07		130	1.99873E-05		130	4.01787E-06		130	4.99302E-06
135	2.1559E-06		135	1.99808E-05		135	6.76739E-07		135	1.99957E-05		135	4.04674E-06		135	3.9182E-06					135	1.99833E-05		135	2.57797E-07		135	5.21239E-07					135	4.14087E-06		135	5.2525E-06
140	2.31315E-06		140	1.99795E-05		140	7.92167E-07		140	1.99778E-05		140	4.92167E-06		140	4.00147E-06					140	1.99838E-05		140	2.60512E-07		140	5.33661E-07					140	4.3011E-06		140	5.48167E-06
145	2.60799E-06		145	1.99795E-05		145	7.60425E-07		145	1.99835E-05		145	6.10786E-06		145	4.09598E-06					145	2.62959E-07		145	5.46828E-07					145	4.49785E-06		145	5.74321E-06
150	2.90717E-06					150	8.35266E-07					150	7.53985E-06		150	4.21064E-06					150	2.65571E-07		150	5.59411E-07					150	4.75347E-06		150	6.03823E-06
155	3.19282E-06					155	8.71465E-07					155	9.77238E-06		155	4.36291E-06					155	2.66544E-07		155	5.70859E-07					155	5.08315E-06		155	6.28938E-06
160	3.52844E-06					160	1.01929E-06					160	1.30668E-05		160	4.55675E-06					160	2.69049E-07		160	5.86094E-07					160	5.44997E-06		160	6.6133E-06
165	3.86926E-06					165	1.0889E-06					165	1.80821E-05		165	4.86468E-06					165	2.71616E-07		165	5.94765E-07					165	5.94482E-06		165	6.87866E-06
170	4.27753E-06					170	1.16814E-06					170	1.9982E-05		170	5.2487E-06					170	2.73683E-07		170	6.14281E-07					170	6.52805E-06		170	7.18553E-06
175	4.64368E-06					175	1.33523E-06					175	1.99812E-05		175	5.86828E-06					175	2.76126E-07		175	6.21327E-07					175	7.19911E-06		175	7.54572E-06
180	5.01252E-06					180	1.50888E-06					180	1.9981E-05		180	6.89832E-06					180	2.78733E-07		180	6.36343E-07					180	7.97104E-06		180	7.87112E-06
185	5.22655E-06					185	1.51093E-06					185	1.99806E-05		185	8.89868E-06					185	2.7985E-07		185	6.48988E-07					185	8.9367E-06		185	8.29011E-06
190	5.56989E-06					190	1.6307E-06					190	1.19636E-05					190	2.83783E-07		190	6.61151E-07					190	9.99745E-06		190	8.69034E-06
195	5.93569E-06					195	1.80827E-06					195	1.55275E-05					195	2.85136E-07		195	6.76551E-07					195	1.11939E-05		195	9.14796E-06
200	6.30896E-06					200	1.9137E-06					200	1.8914E-05					200	2.88539E-07		200	6.9313E-07					200	1.25748E-05		200	9.68038E-06
205	6.71867E-06					205	2.37165E-06					205	1.99763E-05					205	2.90517E-07		205	7.1024E-07					205	1.41809E-05		205	1.02464E-05
210	7.11939E-06					210	2.30878E-06					210	1.99827E-05					210	2.93428E-07		210	7.23693E-07					210	1.58831E-05		210	1.08602E-05
215	7.45726E-06					215	2.49999E-06					215	1.99814E-05					215	2.95866E-07		215	7.39696E-07					215	1.77633E-05		215	1.14709E-05
220	7.91346E-06					220	2.51127E-06					220	1.9978E-05					220	2.99103E-07		220	7.60377E-07					220	1.98397E-05		220	1.243E-05
225	8.37395E-06					225	2.64122E-06					225	3.01411E-07		225	7.82242E-07					225	1.99813E-05		225	1.32902E-05
230	8.77338E-06					230	3.0018E-06					230	3.0405E-07		230	7.89788E-07					230	1.99826E-05		230	1.4229E-05
235	9.20838E-06					235	3.21515E-06					235	3.07265E-07		235	8.12113E-07					235	1.99828E-05		235	1.51397E-05
240	9.71234E-06					240	1.38598E-05					240	3.10319E-07		240	8.43847E-07					240	1.9982E-05		240	1.62201E-05
245	1.02123E-05					245	1.89821E-05					245	3.14311E-07		245	8.6177E-07					245	1.81596E-05
250	1.05766E-05					250	2.00067E-05					250	3.17776E-07		250	8.78371E-07					250	1.93714E-05
255	1.10133E-05					255	1.99543E-05					255	3.20999E-07		255	9.07336E-07					255	1.99821E-05
260	1.1482E-05					260	1.99847E-05					260	3.24795E-07		260	9.20399E-07					260	1.9982E-05
265	1.18603E-05					265	2.00047E-05					265	3.29823E-07		265	9.45756E-07					265	1.99822E-05
270	1.22789E-05					270	3.34122E-07		270	9.76699E-07					270	1.99833E-05
275	1.26782E-05					275	3.38864E-07		275	1.00121E-06
280	1.31286E-05					280	3.43517E-07		280	1.0452E-06
285	1.35167E-05					285	3.50089E-07		285	1.07254E-06
290	1.39708E-05					290	3.5634E-07		290	1.11493E-06
295	1.44543E-05					295	3.59372E-07		295	1.13617E-06
300	1.48618E-05					300	3.6815E-07		300	1.17522E-06
305	1.52391E-05					305	3.73982E-07		305	1.20974E-06
310	1.56817E-05					310	3.79693E-07		310	1.23009E-06
315	1.61439E-05					315	3.91253E-07		315	1.26973E-06
320	1.66616E-05					320	3.95593E-07		320	1.31705E-06
325	1.72394E-05					325	4.06438E-07		325	1.3336E-06
330	1.76965E-05					330	4.14502E-07		330	1.37809E-06
335	1.81364E-05					335	4.23544E-07		335	1.38064E-06
340	1.86465E-05					340	4.36077E-07		340	1.43056E-06
345	1.90532E-05					345	4.46739E-07		345	1.48148E-06
350	1.963E-05					350	4.57424E-07		350	1.52255E-06
355	1.99778E-05					355	4.66466E-07		355	1.56441E-06
360	1.99793E-05					360	4.84559E-07		360	1.61639E-06
365	1.99795E-05					365	5.27638E-07		365	1.67137E-06
370	1.99792E-05					370	6.25074E-07		370	1.7192E-06
					375	1.06586E-06		375	1.76718E-06
					380	2.60309E-06		380	1.81716E-06
					385	5.62517E-06		385	1.8918E-06
					390	9.87786E-06		390	1.95297E-06
					395	1.50275E-05		395	2.03031E-06
					400	1.99807E-05		400	2.11292E-06
					405	1.99835E-05		405	2.17669E-06
					410	1.99832E-05		410	2.29776E-06
					415	1.99838E-05		415	2.39604E-06
					420	2.49314E-06
					425	2.5855E-06
					430	2.76615E-06
					435	2.91764E-06
					440	3.05214E-06
					445	3.20349E-06
					450	3.38254E-06
					455	3.70643E-06
					460	4.02644E-06
					465	4.25758E-06
					470	4.54937E-06
					475	4.8833E-06
					480	5.48042E-06
					485	5.95282E-06
					490	6.35305E-06
					495	7.11612E-06

## IV Curve Plots
20190227 Sensors (UCSC)
IV Curve
Bias (Voltage)
Leakage Current (A)
20190227 Sensors (UCSC)
IV Curve
Bias (Voltage)
Leakage Current (A)

## CNM Data
(CNM) W09-S5					(CNM) W09-S2					(CNM) W07-S6					(CNM) W07-S5					(CNM) W07-S3					(CNM) W07-S2					(CNM) W07-S1					(CNM) W04-S6					(CNM) W04-S3					(CNM) W03-S6					(CNM) W02-S4					(CNM) W02-S1
V	I				V	I				V	I				V	I				V	I				V	I				V	I				V	I				V	I				V	I				V	I				V	I
0	-1.50224E-11	0	1.50224E-11		0	-3.44668E-11	0	3.44668E-11		0	9.47948E-10	0	-9.47948E-10		0	1.13327E-09	0	-1.13327E-09		0	1.90403E-10	0	-1.90403E-10		0	7.38237E-10	0	-7.38237E-10		0	1.18928E-10	0	-1.18928E-10		0	-6.45726E-11	0	6.45726E-11		0	-7.02751E-11	0	7.02751E-11		0	-7.19695E-11	0	7.19695E-11		0	-8.24045E-11	0	8.24045E-11		0	-9.56299E-11	0	9.56299E-11
-5	-1.42677E-07	5	1.42677E-07		-5	-8.5765E-08	5	8.5765E-08		-5	-1.28185E-06	5	1.28185E-06		-5	-1.2779E-06	5	1.2779E-06		-5	-9.27187E-07	5	9.27187E-07		-5	-3.81207E-06	5	3.81207E-06		-5	-9.34121E-07	5	9.34121E-07		-5	-8.53126E-08	5	8.53126E-08		-5	-5.6768E-08	5	5.6768E-08		-5	-9.28911E-08	5	9.28911E-08		-5	-9.49517E-08	5	9.49517E-08		-5	-9.65463E-08	5	9.65463E-08
-10	-1.82919E-07	10	1.82919E-07		-10	-1.18477E-07	10	1.18477E-07		-10	-2.27824E-06	10	2.27824E-06		-10	-2.16627E-06	10	2.16627E-06		-10	-1.64636E-06	10	1.64636E-06		-10	-5.26808E-06	10	5.26808E-06		-10	-1.63841E-06	10	1.63841E-06		-10	-1.50479E-07	10	1.50479E-07		-10	-9.13394E-08	10	9.13394E-08		-10	-1.32043E-07	10	1.32043E-07		-10	-1.42493E-07	10	1.42493E-07		-10	-1.43285E-07	10	1.43285E-07
-15	-2.12548E-07	15	2.12548E-07		-15	-1.46053E-07	15	1.46053E-07		-15	-2.91174E-06	15	2.91174E-06		-15	-2.70726E-06	15	2.70726E-06		-15	-2.11834E-06	15	2.11834E-06		-15	-5.98697E-06	15	5.98697E-06		-15	-2.11642E-06	15	2.11642E-06		-15	-1.97864E-07	15	1.97864E-07		-15	-1.21537E-07	15	1.21537E-07		-15	-1.62206E-07	15	1.62206E-07		-15	-1.82759E-07	15	1.82759E-07		-15	-1.81542E-07	15	1.81542E-07
-20	-2.33674E-07	20	2.33674E-07		-20	-1.65651E-07	20	1.65651E-07		-20	-3.08852E-06	20	3.08852E-06		-20	-2.86053E-06	20	2.86053E-06		-20	-2.24747E-06	20	2.24747E-06		-20	-6.24367E-06	20	6.24367E-06		-20	-2.27323E-06	20	2.27323E-06		-20	-2.11099E-07	20	2.11099E-07		-20	-1.38653E-07	20	1.38653E-07		-20	-1.76562E-07	20	1.76562E-07		-20	-2.09136E-07	20	2.09136E-07		-20	-2.0293E-07	20	2.0293E-07
-25	-2.48014E-07	25	2.48014E-07		-25	-1.76992E-07	25	1.76992E-07		-25	-3.12006E-06	25	3.12006E-06		-25	-2.89011E-06	25	2.89011E-06		-25	-2.27284E-06	25	2.27284E-06		-25	-6.36526E-06	25	6.36526E-06		-25	-2.29506E-06	25	2.29506E-06		-25	-2.21834E-07	25	2.21834E-07		-25	-1.46252E-07	25	1.46252E-07		-25	-1.82358E-07	25	1.82358E-07		-25	-2.22785E-07	25	2.22785E-07		-25	-2.14701E-07	25	2.14701E-07
-30	-2.56243E-07	30	2.56243E-07		-30	-1.8291E-07	30	1.8291E-07		-30	-3.14293E-06	30	3.14293E-06		-30	-2.91204E-06	30	2.91204E-06		-30	-2.29224E-06	30	2.29224E-06		-30	-6.51876E-06	30	6.51876E-06		-30	-2.31017E-06	30	2.31017E-06		-30	-2.37961E-07	30	2.37961E-07		-30	-1.49912E-07	30	1.49912E-07		-30	-1.8549E-07	30	1.8549E-07		-30	-2.29684E-07	30	2.29684E-07		-30	-2.25551E-07	30	2.25551E-07
-35	-2.61807E-07	35	2.61807E-07		-35	-1.86512E-07	35	1.86512E-07		-35	-3.16264E-06	35	3.16264E-06		-35	-2.93128E-06	35	2.93128E-06		-35	-2.30912E-06	35	2.30912E-06		-35	-6.73805E-06	35	6.73805E-06		-35	-2.32285E-06	35	2.32285E-06		-35	-2.55785E-07	35	2.55785E-07		-35	-1.52642E-07	35	1.52642E-07		-35	-1.88439E-07	35	1.88439E-07		-35	-2.31147E-07	35	2.31147E-07		-35	-2.39325E-07	35	2.39325E-07
-40	-2.66631E-07	40	2.66631E-07		-40	-1.89116E-07	40	1.89116E-07		-40	-3.18027E-06	40	3.18027E-06		-40	-2.94875E-06	40	2.94875E-06		-40	-2.32501E-06	40	2.32501E-06		-40	-7.06303E-06	40	7.06303E-06		-40	-2.33487E-06	40	2.33487E-06		-40	-2.74053E-07	40	2.74053E-07		-40	-1.55128E-07	40	1.55128E-07		-40	-1.91165E-07	40	1.91165E-07		-40	-2.32434E-07	40	2.32434E-07		-40	-2.58573E-07	40	2.58573E-07
-45	-2.70173E-07	45	2.70173E-07		-45	-1.90947E-07	45	1.90947E-07		-45	-3.19669E-06	45	3.19669E-06		-45	-2.96517E-06	45	2.96517E-06		-45	-2.34013E-06	45	2.34013E-06		-45	-7.56902E-06	45	7.56902E-06		-45	-2.34587E-06	45	2.34587E-06		-45	-2.91894E-07	45	2.91894E-07		-45	-1.57918E-07	45	1.57918E-07		-45	-1.93795E-07	45	1.93795E-07		-45	-2.33455E-07	45	2.33455E-07		-45	-2.8423E-07	45	2.8423E-07
-50	-2.74933E-07	50	2.74933E-07		-50	-1.92879E-07	50	1.92879E-07		-50	-3.21184E-06	50	3.21184E-06		-50	-2.98107E-06	50	2.98107E-06		-50	-2.35514E-06	50	2.35514E-06		-50	-8.28047E-06	50	8.28047E-06		-50	-2.35677E-06	50	2.35677E-06		-50	-3.11524E-07	50	3.11524E-07		-50	-1.60974E-07	50	1.60974E-07		-50	-1.96781E-07	50	1.96781E-07		-50	-2.35012E-07	50	2.35012E-07		-50	-3.17045E-07	50	3.17045E-07
-55	-2.78551E-07	55	2.78551E-07		-55	-1.95573E-07	55	1.95573E-07		-55	-3.22561E-06	55	3.22561E-06		-55	-2.99647E-06	55	2.99647E-06		-55	-2.37097E-06	55	2.37097E-06		-55	-9.2528E-06	55	9.2528E-06		-55	-2.36774E-06	55	2.36774E-06		-55	-3.30353E-07	55	3.30353E-07		-55	-1.63772E-07	55	1.63772E-07		-55	-1.99508E-07	55	1.99508E-07		-55	-2.37539E-07	55	2.37539E-07		-55	-3.56509E-07	55	3.56509E-07
-60	-2.82957E-07	60	2.82957E-07		-60	-1.97139E-07	60	1.97139E-07		-60	-3.24E-06	60	3.24E-06		-60	-3.01049E-06	60	3.01049E-06		-60	-2.38566E-06	60	2.38566E-06		-60	-1.05088E-05	60	1.05088E-05		-60	-2.37813E-06	60	2.37813E-06		-60	-3.47934E-07	60	3.47934E-07		-60	-1.66663E-07	60	1.66663E-07		-60	-2.02494E-07	60	2.02494E-07		-60	-2.40358E-07	60	2.40358E-07		-60	-4.01375E-07	60	4.01375E-07
-65	-2.86641E-07	65	2.86641E-07		-65	-1.9996E-07	65	1.9996E-07		-65	-3.25318E-06	65	3.25318E-06		-65	-3.02568E-06	65	3.02568E-06		-65	-2.39884E-06	65	2.39884E-06		-65	-1.20328E-05	65	1.20328E-05		-65	-2.38776E-06	65	2.38776E-06		-65	-3.67895E-07	65	3.67895E-07		-65	-1.69375E-07	65	1.69375E-07		-65	-2.05766E-07	65	2.05766E-07		-65	-2.43064E-07	65	2.43064E-07		-65	-4.51541E-07	65	4.51541E-07
-70	-2.90098E-07	70	2.90098E-07		-70	-2.02409E-07	70	2.02409E-07		-70	-3.2677E-06	70	3.2677E-06		-70	-3.03949E-06	70	3.03949E-06		-70	-2.41278E-06	70	2.41278E-06		-70	-1.37654E-05	70	1.37654E-05		-70	-2.39749E-06	70	2.39749E-06		-70	-3.87233E-07	70	3.87233E-07		-70	-1.72885E-07	70	1.72885E-07		-70	-2.0903E-07	70	2.0903E-07		-70	-2.45928E-07	70	2.45928E-07		-70	-5.06083E-07	70	5.06083E-07
-75	-2.94239E-07	75	2.94239E-07		-75	-2.04413E-07	75	2.04413E-07		-75	-3.28119E-06	75	3.28119E-06		-75	-3.05471E-06	75	3.05471E-06		-75	-2.42712E-06	75	2.42712E-06		-75	-1.56968E-05	75	1.56968E-05		-75	-2.40751E-06	75	2.40751E-06		-75	-4.02069E-07	75	4.02069E-07		-75	-1.7561E-07	75	1.7561E-07		-75	-2.12446E-07	75	2.12446E-07		-75	-2.48816E-07	75	2.48816E-07		-75	-5.65631E-07	75	5.65631E-07
-80	-2.98208E-07	80	2.98208E-07		-80	-2.07123E-07	80	2.07123E-07		-80	-3.29707E-06	80	3.29707E-06		-80	-3.06843E-06	80	3.06843E-06		-80	-2.44321E-06	80	2.44321E-06		-80	-1.77783E-05	80	1.77783E-05		-80	-2.4176E-06	80	2.4176E-06		-80	-4.23599E-07	80	4.23599E-07		-80	-1.78428E-07	80	1.78428E-07		-80	-2.15869E-07	80	2.15869E-07		-80	-2.52137E-07	80	2.52137E-07		-80	-6.28318E-07	80	6.28318E-07
-85	-3.01991E-07	85	3.01991E-07		-85	-2.09145E-07	85	2.09145E-07		-85	-3.31369E-06	85	3.31369E-06		-85	-3.08352E-06	85	3.08352E-06		-85	-2.4591E-06	85	2.4591E-06		-85	-2.00561E-05	85	2.00561E-05		-85	-2.42625E-06	85	2.42625E-06		-85	-4.38211E-07	85	4.38211E-07		-85	-1.81367E-07	85	1.81367E-07		-85	-2.19718E-07	85	2.19718E-07		-85	-2.55362E-07	85	2.55362E-07		-85	-6.94081E-07	85	6.94081E-07
-90	-3.05239E-07	90	3.05239E-07		-90	-2.1095E-07	90	2.1095E-07		-90	-3.33344E-06	90	3.33344E-06		-90	-3.09756E-06	90	3.09756E-06		-90	-2.47561E-06	90	2.47561E-06		-90	-2.24067E-05	90	2.24067E-05		-90	-2.4364E-06	90	2.4364E-06		-90	-4.55747E-07	90	4.55747E-07		-90	-1.83215E-07	90	1.83215E-07		-90	-2.23343E-07	90	2.23343E-07		-90	-2.58356E-07	90	2.58356E-07		-90	-7.62472E-07	90	7.62472E-07
-95	-3.09818E-07	95	3.09818E-07		-95	-2.13497E-07	95	2.13497E-07		-95	-3.35539E-06	95	3.35539E-06		-95	-3.1127E-06	95	3.1127E-06		-95	-2.49205E-06	95	2.49205E-06		-95	-2.48627E-05	95	2.48627E-05		-95	-2.44514E-06	95	2.44514E-06		-95	-4.75668E-07	95	4.75668E-07		-95	-1.85819E-07	95	1.85819E-07		-95	-2.26529E-07	95	2.26529E-07		-95	-2.6204E-07	95	2.6204E-07		-95	-8.37756E-07	95	8.37756E-07
-100	-3.15784E-07	100	3.15784E-07		-100	-2.15321E-07	100	2.15321E-07		-100	-3.3781E-06	100	3.3781E-06		-100	-3.12685E-06	100	3.12685E-06		-100	-2.50962E-06	100	2.50962E-06		-100	-2.7434E-05	100	2.7434E-05		-100	-2.45402E-06	100	2.45402E-06		-100	-4.92119E-07	100	4.92119E-07		-100	-1.87846E-07	100	1.87846E-07		-100	-2.29341E-07	100	2.29341E-07		-100	-2.65651E-07	100	2.65651E-07		-100	-9.08929E-07	100	9.08929E-07
-105	-3.23395E-07	105	3.23395E-07		-105	-2.16736E-07	105	2.16736E-07		-105	-3.40294E-06	105	3.40294E-06		-105	-3.14156E-06	105	3.14156E-06		-105	-2.52986E-06	105	2.52986E-06		-105	-3.01687E-05	105	3.01687E-05		-105	-2.46196E-06	105	2.46196E-06		-105	-5.07517E-07	105	5.07517E-07		-105	-1.89619E-07	105	1.89619E-07		-105	-2.32405E-07	105	2.32405E-07		-105	-2.68938E-07	105	2.68938E-07		-105	-9.83178E-07	105	9.83178E-07
-110	-3.35245E-07	110	3.35245E-07		-110	-2.19021E-07	110	2.19021E-07		-110	-3.42887E-06	110	3.42887E-06		-110	-3.15616E-06	110	3.15616E-06		-110	-2.5486E-06	110	2.5486E-06		-110	-3.31877E-05	110	3.31877E-05		-110	-2.47103E-06	110	2.47103E-06		-110	-5.24995E-07	110	5.24995E-07		-110	-1.9226E-07	110	1.9226E-07		-110	-2.35165E-07	110	2.35165E-07		-110	-2.72829E-07	110	2.72829E-07		-110	-1.05907E-06	110	1.05907E-06
-115	-3.55791E-07	115	3.55791E-07		-115	-2.20756E-07	115	2.20756E-07		-115	-3.46258E-06	115	3.46258E-06		-115	-3.17093E-06	115	3.17093E-06		-115	-2.56825E-06	115	2.56825E-06		-115	-3.64825E-05	115	3.64825E-05		-115	-2.47935E-06	115	2.47935E-06		-115	-5.4305E-07	115	5.4305E-07		-115	-1.93791E-07	115	1.93791E-07		-115	-2.37275E-07	115	2.37275E-07		-115	-2.7721E-07	115	2.7721E-07		-115	-1.13342E-06	115	1.13342E-06
-120	-3.88265E-07	120	3.88265E-07		-120	-2.22721E-07	120	2.22721E-07		-120	-3.50212E-06	120	3.50212E-06		-120	-3.18595E-06	120	3.18595E-06		-120	-2.58901E-06	120	2.58901E-06		-120	-4.00025E-05	120	4.00025E-05		-120	-2.48773E-06	120	2.48773E-06		-120	-5.6227E-07	120	5.6227E-07		-120	-1.96031E-07	120	1.96031E-07		-120	-2.39754E-07	120	2.39754E-07		-120	-2.82315E-07	120	2.82315E-07		-120	-1.21422E-06	120	1.21422E-06
-125	-4.34021E-07	125	4.34021E-07		-125	-2.25116E-07	125	2.25116E-07		-125	-3.54886E-06	125	3.54886E-06		-125	-3.20065E-06	125	3.20065E-06		-125	-2.61536E-06	125	2.61536E-06		-125	-4.37106E-05	125	4.37106E-05		-125	-2.49699E-06	125	2.49699E-06		-125	-5.79546E-07	125	5.79546E-07		-125	-1.9863E-07	125	1.9863E-07		-125	-2.42604E-07	125	2.42604E-07		-125	-2.88293E-07	125	2.88293E-07		-125	-1.2948E-06	125	1.2948E-06
-130	-4.94932E-07	130	4.94932E-07		-130	-2.26641E-07	130	2.26641E-07		-130	-3.60915E-06	130	3.60915E-06		-130	-3.21625E-06	130	3.21625E-06		-130	-2.64058E-06	130	2.64058E-06		-130	-4.76459E-05	130	4.76459E-05		-130	-2.50544E-06	130	2.50544E-06		-130	-6.0259E-07	130	6.0259E-07		-130	-2.00305E-07	130	2.00305E-07		-130	-2.4512E-07	130	2.4512E-07		-130	-2.95632E-07	130	2.95632E-07		-130	-1.37221E-06	130	1.37221E-06
-135	-5.73333E-07	135	5.73333E-07		-135	-2.28873E-07	135	2.28873E-07		-135	-3.67921E-06	135	3.67921E-06		-135	-3.23201E-06	135	3.23201E-06		-135	-2.67378E-06	135	2.67378E-06		-135	-5.21002E-05	135	5.21002E-05		-135	-2.51521E-06	135	2.51521E-06		-135	-6.17239E-07	135	6.17239E-07		-135	-2.03476E-07	135	2.03476E-07		-135	-2.47685E-07	135	2.47685E-07		-135	-3.0452E-07	135	3.0452E-07		-135	-1.45463E-06	135	1.45463E-06
-140	-6.69354E-07	140	6.69354E-07		-140	-2.31643E-07	140	2.31643E-07		-140	-3.77186E-06	140	3.77186E-06		-140	-3.24764E-06	140	3.24764E-06		-140	-2.70786E-06	140	2.70786E-06		-140	-5.70449E-05	140	5.70449E-05		-140	-2.525E-06	140	2.525E-06		-140	-6.41389E-07	140	6.41389E-07		-140	-2.05696E-07	140	2.05696E-07		-140	-2.5063E-07	140	2.5063E-07		-140	-3.16107E-07	140	3.16107E-07		-140	-1.52783E-06	140	1.52783E-06
-145	-7.79132E-07	145	7.79132E-07		-145	-2.33786E-07	145	2.33786E-07		-145	-3.88563E-06	145	3.88563E-06		-145	-3.26289E-06	145	3.26289E-06		-145	-2.74812E-06	145	2.74812E-06		-145	-6.24704E-05	145	6.24704E-05		-145	-2.53252E-06	145	2.53252E-06		-145	-6.6325E-07	145	6.6325E-07		-145	-2.0854E-07	145	2.0854E-07		-145	-2.53641E-07	145	2.53641E-07		-145	-3.30025E-07	145	3.30025E-07		-145	-1.59961E-06	145	1.59961E-06
-150	-9.05197E-07	150	9.05197E-07		-150	-2.36822E-07	150	2.36822E-07		-150	-4.0238E-06	150	4.0238E-06		-150	-3.27876E-06	150	3.27876E-06		-150	-2.79249E-06	150	2.79249E-06		-150	-6.85393E-05	150	6.85393E-05		-150	-2.54338E-06	150	2.54338E-06		-150	-6.85826E-07	150	6.85826E-07		-150	-2.1182E-07	150	2.1182E-07		-150	-2.56585E-07	150	2.56585E-07		-150	-3.48064E-07	150	3.48064E-07		-150	-1.67354E-06	150	1.67354E-06
-155	-1.04572E-06	155	1.04572E-06		-155	-2.41191E-07	155	2.41191E-07		-155	-4.20135E-06	155	4.20135E-06		-155	-3.29407E-06	155	3.29407E-06		-155	-2.84175E-06	155	2.84175E-06		-155	-7.49079E-05	155	7.49079E-05		-155	-2.55275E-06	155	2.55275E-06		-155	-7.08541E-07	155	7.08541E-07		-155	-2.14447E-07	155	2.14447E-07		-155	-2.60242E-07	155	2.60242E-07		-155	-3.69987E-07	155	3.69987E-07		-155	-1.75727E-06	155	1.75727E-06
-160	-1.20258E-06	160	1.20258E-06		-160	-2.42738E-07	160	2.42738E-07		-160	-4.41028E-06	160	4.41028E-06		-160	-3.31038E-06	160	3.31038E-06		-160	-2.88887E-06	160	2.88887E-06		-160	-8.14602E-05	160	8.14602E-05		-160	-2.56288E-06	160	2.56288E-06		-160	-7.29766E-07	160	7.29766E-07		-160	-2.17413E-07	160	2.17413E-07		-160	-2.63858E-07	160	2.63858E-07		-160	-3.97026E-07	160	3.97026E-07		-160	-1.83291E-06	160	1.83291E-06
-165	-1.37155E-06	165	1.37155E-06		-165	-2.46287E-07	165	2.46287E-07		-165	-4.68208E-06	165	4.68208E-06		-165	-3.32471E-06	165	3.32471E-06		-165	-2.94823E-06	165	2.94823E-06		-165	-8.82378E-05	165	8.82378E-05		-165	-2.57259E-06	165	2.57259E-06		-165	-7.55059E-07	165	7.55059E-07		-165	-2.2098E-07	165	2.2098E-07		-165	-2.67214E-07	165	2.67214E-07		-165	-4.29171E-07	165	4.29171E-07		-165	-1.91706E-06	165	1.91706E-06
-170	-1.55745E-06	170	1.55745E-06		-170	-2.4965E-07	170	2.4965E-07		-170	-4.98946E-06	170	4.98946E-06		-170	-3.34111E-06	170	3.34111E-06		-170	-2.99868E-06	170	2.99868E-06		-170	-9.50992E-05	170	9.50992E-05		-170	-2.58209E-06	170	2.58209E-06		-170	-7.84363E-07	170	7.84363E-07		-170	-2.23841E-07	170	2.23841E-07		-170	-2.7124E-07	170	2.7124E-07		-170	-4.68089E-07	170	4.68089E-07		-170	-2.00123E-06	170	2.00123E-06
-175	-1.75577E-06	175	1.75577E-06		-175	-2.53026E-07	175	2.53026E-07		-175	-5.408E-06	175	5.408E-06		-175	-3.35664E-06	175	3.35664E-06		-175	-3.05387E-06	175	3.05387E-06		-175	-9.99996E-05	175	9.99996E-05		-175	-2.59256E-06	175	2.59256E-06		-175	-8.14184E-07	175	8.14184E-07		-175	-2.27338E-07	175	2.27338E-07		-175	-2.75105E-07	175	2.75105E-07		-175	-5.13492E-07	175	5.13492E-07		-175	-2.09031E-06	175	2.09031E-06
-180	-1.97076E-06	180	1.97076E-06		-180	-2.58227E-07	180	2.58227E-07		-180	-5.8382E-06	180	5.8382E-06		-180	-3.37392E-06	180	3.37392E-06		-180	-3.11439E-06	180	3.11439E-06		-180	-9.99997E-05	180	9.99997E-05		-180	-2.60283E-06	180	2.60283E-06		-180	-8.41029E-07	180	8.41029E-07		-180	-2.31034E-07	180	2.31034E-07		-180	-2.79007E-07	180	2.79007E-07		-180	-5.6524E-07	180	5.6524E-07		-180	-2.18542E-06	180	2.18542E-06
-185	-2.20164E-06	185	2.20164E-06		-185	-2.61483E-07	185	2.61483E-07		-185	-6.37486E-06	185	6.37486E-06		-185	-3.3893E-06	185	3.3893E-06		-185	-3.19786E-06	185	3.19786E-06		-185	-9.99996E-05	185	9.99996E-05		-185	-2.61399E-06	185	2.61399E-06		-185	-8.70756E-07	185	8.70756E-07		-185	-2.34044E-07	185	2.34044E-07		-185	-2.83506E-07	185	2.83506E-07		-185	-6.23444E-07	185	6.23444E-07		-185	-2.27321E-06	185	2.27321E-06
-190	-2.44824E-06	190	2.44824E-06		-190	-2.65734E-07	190	2.65734E-07		-190	-7.02018E-06	190	7.02018E-06		-190	-3.40549E-06	190	3.40549E-06		-190	-3.26294E-06	190	3.26294E-06		-190	-9.99996E-05	190	9.99996E-05		-190	-2.62456E-06	190	2.62456E-06		-190	-9.01833E-07	190	9.01833E-07		-190	-2.37856E-07	190	2.37856E-07		-190	-2.877E-07	190	2.877E-07		-190	-6.90157E-07	190	6.90157E-07		-190	-2.40118E-06	190	2.40118E-06
-195	-2.70944E-06	195	2.70944E-06		-195	-2.71381E-07	195	2.71381E-07		-195	-7.72151E-06	195	7.72151E-06		-195	-3.42273E-06	195	3.42273E-06		-195	-3.34168E-06	195	3.34168E-06		-195	-9.99996E-05	195	9.99996E-05		-195	-2.63486E-06	195	2.63486E-06		-195	-9.30569E-07	195	9.30569E-07		-195	-2.41755E-07	195	2.41755E-07		-195	-2.92139E-07	195	2.92139E-07		-195	-7.63346E-07	195	7.63346E-07		-195	-2.54544E-06	195	2.54544E-06
-200	-2.98878E-06	200	2.98878E-06		-200	-2.78132E-07	200	2.78132E-07		-200	-8.56076E-06	200	8.56076E-06		-200	-3.44027E-06	200	3.44027E-06		-200	-3.42838E-06	200	3.42838E-06		-200	-9.99997E-05	200	9.99997E-05		-200	-2.64317E-06	200	2.64317E-06		-200	-9.65416E-07	200	9.65416E-07		-200	-2.44708E-07	200	2.44708E-07		-200	-2.97057E-07	200	2.97057E-07		-200	-8.43964E-07	200	8.43964E-07		-200	-2.707E-06	200	2.707E-06
-205	-3.2873E-06	205	3.2873E-06		-205	-2.84825E-07	205	2.84825E-07		-205	-9.49712E-06	205	9.49712E-06		-205	-3.4573E-06	205	3.4573E-06		-205	-3.49525E-06	205	3.49525E-06		-205	-9.99996E-05	205	9.99996E-05		-205	-2.65526E-06	205	2.65526E-06		-205	-9.98362E-07	205	9.98362E-07		-205	-2.48988E-07	205	2.48988E-07		-205	-3.01766E-07	205	3.01766E-07		-205	-9.3194E-07	205	9.3194E-07		-205	-2.86807E-06	205	2.86807E-06
-210	-3.60657E-06	210	3.60657E-06		-210	-2.92704E-07	210	2.92704E-07		-210	-1.05249E-05	210	1.05249E-05		-210	-3.47432E-06	210	3.47432E-06		-210	-3.59757E-06	210	3.59757E-06		-210	-9.99996E-05	210	9.99996E-05		-210	-2.66316E-06	210	2.66316E-06		-210	-1.03513E-06	210	1.03513E-06		-210	-2.53294E-07	210	2.53294E-07		-210	-3.06967E-07	210	3.06967E-07		-210	-1.02742E-06	210	1.02742E-06		-210	-3.05003E-06	210	3.05003E-06
-215	-3.93373E-06	215	3.93373E-06		-215	-3.01644E-07	215	3.01644E-07		-215	-1.17593E-05	215	1.17593E-05		-215	-3.49242E-06	215	3.49242E-06		-215	-3.68663E-06	215	3.68663E-06		-215	-9.99996E-05	215	9.99996E-05		-215	-2.67186E-06	215	2.67186E-06		-215	-1.07303E-06	215	1.07303E-06		-215	-2.57155E-07	215	2.57155E-07		-215	-3.1282E-07	215	3.1282E-07		-215	-1.13085E-06	215	1.13085E-06		-215	-3.23038E-06	215	3.23038E-06
-220	-4.2819E-06	220	4.2819E-06		-220	-3.1166E-07	220	3.1166E-07		-220	-1.31127E-05	220	1.31127E-05		-220	-3.50924E-06	220	3.50924E-06		-220	-3.80109E-06	220	3.80109E-06		-220	-9.99996E-05	220	9.99996E-05		-220	-2.68434E-06	220	2.68434E-06		-220	-1.10664E-06	220	1.10664E-06		-220	-2.61673E-07	220	2.61673E-07		-220	-3.18308E-07	220	3.18308E-07		-220	-1.24075E-06	220	1.24075E-06		-220	-3.43238E-06	220	3.43238E-06
-225	-4.66293E-06	225	4.66293E-06		-225	-3.22715E-07	225	3.22715E-07		-225	-1.46372E-05	225	1.46372E-05		-225	-3.52733E-06	225	3.52733E-06		-225	-3.90795E-06	225	3.90795E-06		-225	-9.99996E-05	225	9.99996E-05		-225	-2.694E-06	225	2.694E-06		-225	-1.14861E-06	225	1.14861E-06		-225	-2.66393E-07	225	2.66393E-07		-225	-3.25185E-07	225	3.25185E-07		-225	-1.35934E-06	225	1.35934E-06		-225	-3.63916E-06	225	3.63916E-06
-230	-5.03524E-06	230	5.03524E-06		-230	-3.35036E-07	230	3.35036E-07		-230	-1.6278E-05	230	1.6278E-05		-230	-3.5458E-06	230	3.5458E-06		-230	-3.986E-06	230	3.986E-06		-230	-9.99996E-05	230	9.99996E-05		-230	-2.70632E-06	230	2.70632E-06		-230	-1.20212E-06	230	1.20212E-06		-230	-2.70192E-07	230	2.70192E-07		-230	-3.31009E-07	230	3.31009E-07		-230	-1.48417E-06	230	1.48417E-06		-230	-3.86209E-06	230	3.86209E-06
-235	-5.43876E-06	235	5.43876E-06		-235	-3.48642E-07	235	3.48642E-07		-235	-1.81182E-05	235	1.81182E-05		-235	-3.56678E-06	235	3.56678E-06		-235	-4.10061E-06	235	4.10061E-06		-235	-9.99996E-05	235	9.99996E-05		-235	-2.71563E-06	235	2.71563E-06		-235	-1.24953E-06	235	1.24953E-06		-235	-2.75301E-07	235	2.75301E-07		-235	-3.37783E-07	235	3.37783E-07		-235	-1.61586E-06	235	1.61586E-06		-235	-4.08058E-06	235	4.08058E-06
-240	-5.85322E-06	240	5.85322E-06		-240	-3.64538E-07	240	3.64538E-07		-240	-2.01298E-05	240	2.01298E-05		-240	-3.58563E-06	240	3.58563E-06		-240	-4.17558E-06	240	4.17558E-06		-240	-9.99996E-05	240	9.99996E-05		-240	-2.72553E-06	240	2.72553E-06		-240	-1.29679E-06	240	1.29679E-06		-240	-2.80431E-07	240	2.80431E-07		-240	-3.4537E-07	240	3.4537E-07		-240	-1.75588E-06	240	1.75588E-06		-240	-4.31733E-06	240	4.31733E-06
-245	-6.30014E-06	245	6.30014E-06		-245	-3.8156E-07	245	3.8156E-07		-245	-2.23716E-05	245	2.23716E-05		-245	-3.6064E-06	245	3.6064E-06		-245	-4.25726E-06	245	4.25726E-06		-245	-9.99996E-05	245	9.99996E-05		-245	-2.73506E-06	245	2.73506E-06		-245	-1.3542E-06	245	1.3542E-06		-245	-2.85734E-07	245	2.85734E-07		-245	-3.52193E-07	245	3.52193E-07		-245	-1.9036E-06	245	1.9036E-06		-245	-4.56891E-06	245	4.56891E-06
-250	-6.76117E-06	250	6.76117E-06		-250	-4.01461E-07	250	4.01461E-07		-250	-2.47616E-05	250	2.47616E-05		-250	-3.62739E-06	250	3.62739E-06		-250	-4.32252E-06	250	4.32252E-06		-250	-9.99995E-05	250	9.99995E-05		-250	-2.74844E-06	250	2.74844E-06		-250	-1.42444E-06	250	1.42444E-06		-250	-2.90918E-07	250	2.90918E-07		-250	-3.60995E-07	250	3.60995E-07		-250	-2.05913E-06	250	2.05913E-06		-250	-4.85442E-06	250	4.85442E-06
-255	-7.23078E-06	255	7.23078E-06		-255	-4.23399E-07	255	4.23399E-07		-255	-2.73605E-05	255	2.73605E-05		-255	-3.64946E-06	255	3.64946E-06		-255	-4.37286E-06	255	4.37286E-06		-255	-9.99994E-05	255	9.99994E-05		-255	-2.75956E-06	255	2.75956E-06		-255	-1.48296E-06	255	1.48296E-06		-255	-2.96104E-07	255	2.96104E-07		-255	-3.70675E-07	255	3.70675E-07		-255	-2.22178E-06	255	2.22178E-06		-255	-5.12661E-06	255	5.12661E-06
-260	-7.7401E-06	260	7.7401E-06		-260	-4.48475E-07	260	4.48475E-07		-260	-3.02273E-05	260	3.02273E-05		-260	-3.67106E-06	260	3.67106E-06		-260	-4.40664E-06	260	4.40664E-06		-260	-9.99996E-05	260	9.99996E-05		-260	-2.76826E-06	260	2.76826E-06		-260	-1.5466E-06	260	1.5466E-06		-260	-3.01967E-07	260	3.01967E-07		-260	-3.82194E-07	260	3.82194E-07		-260	-2.39157E-06	260	2.39157E-06		-260	-5.4295E-06	260	5.4295E-06
-265	-8.23826E-06	265	8.23826E-06		-265	-4.75649E-07	265	4.75649E-07		-265	-3.33019E-05	265	3.33019E-05		-265	-3.69341E-06	265	3.69341E-06		-265	-4.47016E-06	265	4.47016E-06		-265	-9.99994E-05	265	9.99994E-05		-265	-2.77587E-06	265	2.77587E-06		-265	-1.6229E-06	265	1.6229E-06		-265	-3.07633E-07	265	3.07633E-07		-265	-3.9822E-07	265	3.9822E-07		-265	-2.56907E-06	265	2.56907E-06		-265	-5.70929E-06	265	5.70929E-06
-270	-8.75789E-06	270	8.75789E-06		-270	-5.06535E-07	270	5.06535E-07		-270	-3.65762E-05	270	3.65762E-05		-270	-3.71564E-06	270	3.71564E-06		-270	-4.53009E-06	270	4.53009E-06		-270	-9.99996E-05	270	9.99996E-05		-270	-2.7927E-06	270	2.7927E-06		-270	-1.72509E-06	270	1.72509E-06		-270	-3.14188E-07	270	3.14188E-07		-270	-4.20446E-07	270	4.20446E-07		-270	-2.75358E-06	270	2.75358E-06		-270	-6.01551E-06	270	6.01551E-06
-275	-9.32584E-06	275	9.32584E-06		-275	-5.41122E-07	275	5.41122E-07		-275	-4.01512E-05	275	4.01512E-05		-275	-3.73966E-06	275	3.73966E-06		-275	-4.58319E-06	275	4.58319E-06		-275	-9.99996E-05	275	9.99996E-05		-275	-2.80469E-06	275	2.80469E-06		-275	-1.81675E-06	275	1.81675E-06		-275	-3.20696E-07	275	3.20696E-07		-275	-4.52949E-07	275	4.52949E-07		-275	-2.94516E-06	275	2.94516E-06		-275	-6.34074E-06	275	6.34074E-06
-280	-9.87928E-06	280	9.87928E-06		-280	-5.78059E-07	280	5.78059E-07		-280	-4.38455E-05	280	4.38455E-05		-280	-3.76414E-06	280	3.76414E-06		-280	-4.63365E-06	280	4.63365E-06		-280	-9.99996E-05	280	9.99996E-05		-280	-2.81372E-06	280	2.81372E-06		-280	-1.89007E-06	280	1.89007E-06		-280	-3.26857E-07	280	3.26857E-07		-280	-5.01411E-07	280	5.01411E-07		-280	-3.14525E-06	280	3.14525E-06		-280	-6.74289E-06	280	6.74289E-06
-285	-1.04616E-05	285	1.04616E-05		-285	-6.22572E-07	285	6.22572E-07		-285	-4.79157E-05	285	4.79157E-05		-285	-3.78832E-06	285	3.78832E-06		-285	-4.70089E-06	285	4.70089E-06		-285	-9.99996E-05	285	9.99996E-05		-285	-2.82359E-06	285	2.82359E-06		-285	-2E-06	285	2E-06		-285	-3.34788E-07	285	3.34788E-07		-285	-5.68843E-07	285	5.68843E-07		-285	-3.35293E-06	285	3.35293E-06		-285	-7.1264E-06	285	7.1264E-06
-290	-1.10419E-05	290	1.10419E-05		-290	-6.67325E-07	290	6.67325E-07		-290	-5.21636E-05	290	5.21636E-05		-290	-3.81445E-06	290	3.81445E-06		-290	-4.77398E-06	290	4.77398E-06		-290	-9.99996E-05	290	9.99996E-05		-290	-2.83522E-06	290	2.83522E-06		-290	-2.08024E-06	290	2.08024E-06		-290	-3.41478E-07	290	3.41478E-07		-290	-6.67307E-07	290	6.67307E-07		-290	-3.57084E-06	290	3.57084E-06		-290	-7.46565E-06	290	7.46565E-06
-295	-1.16343E-05	295	1.16343E-05		-295	-7.16584E-07	295	7.16584E-07		-295	-5.67094E-05	295	5.67094E-05		-295	-3.84112E-06	295	3.84112E-06		-295	-4.81871E-06	295	4.81871E-06		-295	-9.99996E-05	295	9.99996E-05		-295	-2.85015E-06	295	2.85015E-06		-295	-2.16494E-06	295	2.16494E-06		-295	-3.49152E-07	295	3.49152E-07		-295	-7.93538E-07	295	7.93538E-07		-295	-3.79381E-06	295	3.79381E-06		-295	-7.83951E-06	295	7.83951E-06
-300	-1.22289E-05	300	1.22289E-05		-300	-7.71043E-07	300	7.71043E-07		-300	-6.15778E-05	300	6.15778E-05		-300	-3.8697E-06	300	3.8697E-06		-300	-4.92018E-06	300	4.92018E-06		-300	-9.99996E-05	300	9.99996E-05		-300	-2.86198E-06	300	2.86198E-06		-300	-2.32123E-06	300	2.32123E-06		-300	-3.56687E-07	300	3.56687E-07		-300	-9.91493E-07	300	9.91493E-07		-300	-4.02697E-06	300	4.02697E-06		-300	-8.23853E-06	300	8.23853E-06
-305	-1.28761E-05	305	1.28761E-05		-305	-8.30462E-07	305	8.30462E-07		-305	-6.66667E-05	305	6.66667E-05		-305	-3.89744E-06	305	3.89744E-06		-305	-4.99322E-06	305	4.99322E-06		-305	-9.99995E-05	305	9.99995E-05		-305	-2.87297E-06	305	2.87297E-06		-305	-2.34599E-06	305	2.34599E-06		-305	-3.64755E-07	305	3.64755E-07		-305	-1.28272E-06	305	1.28272E-06		-305	-4.26806E-06	305	4.26806E-06		-305	-8.6681E-06	305	8.6681E-06
-310	-1.35071E-05	310	1.35071E-05		-310	-8.9501E-07	310	8.9501E-07		-310	-7.21303E-05	310	7.21303E-05		-310	-3.92465E-06	310	3.92465E-06		-310	-5.06971E-06	310	5.06971E-06		-310	-9.99996E-05	310	9.99996E-05		-310	-2.88635E-06	310	2.88635E-06		-310	-2.49619E-06	310	2.49619E-06		-310	-3.73627E-07	310	3.73627E-07		-310	-1.46905E-06	310	1.46905E-06		-310	-4.5179E-06	310	4.5179E-06		-310	-9.08813E-06	310	9.08813E-06
-315	-1.4103E-05	315	1.4103E-05		-315	-9.64676E-07	315	9.64676E-07		-315	-7.77183E-05	315	7.77183E-05		-315	-3.95296E-06	315	3.95296E-06		-315	-5.18412E-06	315	5.18412E-06		-315	-9.99996E-05	315	9.99996E-05		-315	-2.90162E-06	315	2.90162E-06		-315	-2.59737E-06	315	2.59737E-06		-315	-3.82667E-07	315	3.82667E-07		-315	-1.77657E-06	315	1.77657E-06		-315	-4.77549E-06	315	4.77549E-06		-315	-9.49483E-06	315	9.49483E-06
-320	-1.4759E-05	320	1.4759E-05		-320	-1.04077E-06	320	1.04077E-06		-320	-8.37714E-05	320	8.37714E-05		-320	-3.98621E-06	320	3.98621E-06		-320	-5.27988E-06	320	5.27988E-06		-320	-9.99994E-05	320	9.99994E-05		-320	-2.91127E-06	320	2.91127E-06		-320	-2.7081E-06	320	2.7081E-06		-320	-3.91935E-07	320	3.91935E-07		-320	-2.16597E-06	320	2.16597E-06		-320	-5.04043E-06	320	5.04043E-06		-320	-9.96912E-06	320	9.96912E-06
-325	-1.5447E-05	325	1.5447E-05		-325	-1.12417E-06	325	1.12417E-06		-325	-8.99546E-05	325	8.99546E-05		-325	-4.01924E-06	325	4.01924E-06		-325	-5.38726E-06	325	5.38726E-06		-325	-9.99996E-05	325	9.99996E-05		-325	-2.9244E-06	325	2.9244E-06		-325	-2.82387E-06	325	2.82387E-06		-325	-4.01122E-07	325	4.01122E-07		-325	-2.60211E-06	325	2.60211E-06		-325	-5.31449E-06	325	5.31449E-06		-325	-1.04749E-05	325	1.04749E-05
-330	-1.61201E-05	330	1.61201E-05		-330	-1.20343E-06	330	1.20343E-06		-330	-9.66405E-05	330	9.66405E-05		-330	-4.05361E-06	330	4.05361E-06		-330	-5.54252E-06	330	5.54252E-06		-330	-9.99996E-05	330	9.99996E-05		-330	-2.93936E-06	330	2.93936E-06		-330	-2.96109E-06	330	2.96109E-06		-330	-4.11772E-07	330	4.11772E-07		-330	-3.09732E-06	330	3.09732E-06		-330	-5.59661E-06	330	5.59661E-06		-330	-1.09171E-05	330	1.09171E-05
-335	-1.68268E-05	335	1.68268E-05		-335	-1.29043E-06	335	1.29043E-06		-335	-9.99994E-05	335	9.99994E-05		-335	-4.09254E-06	335	4.09254E-06		-335	-5.67966E-06	335	5.67966E-06		-335	-9.99996E-05	335	9.99996E-05		-335	-2.95581E-06	335	2.95581E-06		-335	-3.03101E-06	335	3.03101E-06		-335	-4.21973E-07	335	4.21973E-07		-335	-3.66906E-06	335	3.66906E-06		-335	-5.88653E-06	335	5.88653E-06		-335	-1.13893E-05	335	1.13893E-05
-340	-1.74941E-05	340	1.74941E-05		-340	-1.38831E-06	340	1.38831E-06		-340	-9.99994E-05	340	9.99994E-05		-340	-4.12886E-06	340	4.12886E-06		-340	-5.85119E-06	340	5.85119E-06		-340	-9.99995E-05	340	9.99995E-05		-340	-2.96928E-06	340	2.96928E-06		-340	-3.17018E-06	340	3.17018E-06		-340	-4.34102E-07	340	4.34102E-07		-340	-4.29718E-06	340	4.29718E-06		-340	-6.18709E-06	340	6.18709E-06		-340	-1.19528E-05	340	1.19528E-05
-345	-1.82597E-05	345	1.82597E-05		-345	-1.49441E-06	345	1.49441E-06		-345	-9.99994E-05	345	9.99994E-05		-345	-4.1699E-06	345	4.1699E-06		-345	-6.02496E-06	345	6.02496E-06		-345	-9.99995E-05	345	9.99995E-05		-345	-2.98427E-06	345	2.98427E-06		-345	-3.28843E-06	345	3.28843E-06		-345	-4.57593E-07	345	4.57593E-07		-345	-4.94797E-06	345	4.94797E-06		-345	-6.4952E-06	345	6.4952E-06		-345	-1.25083E-05	345	1.25083E-05
-350	-1.88272E-05	350	1.88272E-05		-350	-1.60429E-06	350	1.60429E-06		-350	-9.99996E-05	350	9.99996E-05		-350	-4.21844E-06	350	4.21844E-06		-350	-6.20255E-06	350	6.20255E-06		-350	-9.99995E-05	350	9.99995E-05		-350	-3.00049E-06	350	3.00049E-06		-350	-3.41206E-06	350	3.41206E-06		-350	-5.17834E-07	350	5.17834E-07		-350	-5.63945E-06	350	5.63945E-06		-350	-6.81071E-06	350	6.81071E-06		-350	-1.30021E-05	350	1.30021E-05
-355	-1.9617E-05	355	1.9617E-05		-355	-1.71818E-06	355	1.71818E-06		-355	-9.99996E-05	355	9.99996E-05		-355	-4.2746E-06	355	4.2746E-06		-355	-6.41339E-06	355	6.41339E-06		-355	-9.99994E-05	355	9.99994E-05		-355	-3.01762E-06	355	3.01762E-06		-355	-3.42855E-06	355	3.42855E-06		-355	-8.18649E-07	355	8.18649E-07		-355	-6.38271E-06	355	6.38271E-06		-355	-7.13493E-06	355	7.13493E-06		-355	-1.35404E-05	355	1.35404E-05
-360	-2.01992E-05	360	2.01992E-05		-360	-1.83849E-06	360	1.83849E-06		-360	-9.99994E-05	360	9.99994E-05		-360	-4.33984E-06	360	4.33984E-06		-360	-6.66754E-06	360	6.66754E-06		-360	-9.99994E-05	360	9.99994E-05		-360	-3.0331E-06	360	3.0331E-06		-360	-3.50588E-06	360	3.50588E-06		-360	-2.05938E-06	360	2.05938E-06		-360	-7.17435E-06	360	7.17435E-06		-360	-7.46644E-06	360	7.46644E-06		-360	-1.41472E-05	360	1.41472E-05
-365	-2.11006E-05	365	2.11006E-05		-365	-1.96495E-06	365	1.96495E-06		-365	-9.99995E-05	365	9.99995E-05		-365	-4.41637E-06	365	4.41637E-06		-365	-6.88949E-06	365	6.88949E-06		-365	-9.99996E-05	365	9.99996E-05		-365	-3.04739E-06	365	3.04739E-06		-365	-3.70744E-06	365	3.70744E-06		-365	-3.6976E-06	365	3.6976E-06		-365	-8.01021E-06	365	8.01021E-06		-365	-7.80887E-06	365	7.80887E-06		-365	-1.46341E-05	365	1.46341E-05
-370	-2.21008E-05	370	2.21008E-05		-370	-2.09987E-06	370	2.09987E-06		-370	-9.99995E-05	370	9.99995E-05		-370	-4.50576E-06	370	4.50576E-06		-370	-7.18064E-06	370	7.18064E-06		-370	-9.99994E-05	370	9.99994E-05		-370	-3.06192E-06	370	3.06192E-06		-370	-3.7759E-06	370	3.7759E-06		-370	-6.30829E-06	370	6.30829E-06		-370	-8.87767E-06	370	8.87767E-06		-370	-8.16026E-06	370	8.16026E-06		-370	-1.5275E-05	370	1.5275E-05
-375	-2.28878E-05	375	2.28878E-05		-375	-2.23994E-06	375	2.23994E-06		-375	-9.99995E-05	375	9.99995E-05		-375	-4.60294E-06	375	4.60294E-06		-375	-7.50416E-06	375	7.50416E-06		-375	-9.99994E-05	375	9.99994E-05		-375	-3.08071E-06	375	3.08071E-06		-375	-3.75188E-06	375	3.75188E-06		-375	-1.01104E-05	375	1.01104E-05		-375	-9.79454E-06	375	9.79454E-06		-375	-8.51982E-06	375	8.51982E-06		-375	-1.58108E-05	375	1.58108E-05
-380	-2.36932E-05	380	2.36932E-05		-380	-2.38663E-06	380	2.38663E-06		-380	-9.99996E-05	380	9.99996E-05		-380	-4.71325E-06	380	4.71325E-06		-380	-7.84586E-06	380	7.84586E-06		-380	-9.99995E-05	380	9.99995E-05		-380	-3.09599E-06	380	3.09599E-06		-380	-3.88929E-06	380	3.88929E-06		-380	-1.56254E-05	380	1.56254E-05		-380	-1.0734E-05	380	1.0734E-05		-380	-8.88733E-06	380	8.88733E-06		-380	-1.65384E-05	380	1.65384E-05
-385	-2.44503E-05	385	2.44503E-05		-385	-2.54312E-06	385	2.54312E-06		-385	-9.99996E-05	385	9.99996E-05		-385	-4.80009E-06	385	4.80009E-06		-385	-8.22898E-06	385	8.22898E-06		-385	-9.99995E-05	385	9.99995E-05		-385	-3.11542E-06	385	3.11542E-06		-385	-3.94615E-06	385	3.94615E-06		-385	-2.22629E-05	385	2.22629E-05		-385	-1.16906E-05	385	1.16906E-05		-385	-9.2653E-06	385	9.2653E-06		-385	-1.71879E-05	385	1.71879E-05
-390	-2.53154E-05	390	2.53154E-05		-390	-2.70441E-06	390	2.70441E-06		-390	-9.99996E-05	390	9.99996E-05		-390	-4.92121E-06	390	4.92121E-06		-390	-8.62966E-06	390	8.62966E-06		-390	-9.99994E-05	390	9.99994E-05		-390	-3.13007E-06	390	3.13007E-06		-390	-4.05453E-06	390	4.05453E-06		-390	-3.11167E-05	390	3.11167E-05		-390	-1.26811E-05	390	1.26811E-05		-390	-9.64643E-06	390	9.64643E-06		-390	-1.76933E-05	390	1.76933E-05
-395	-2.62087E-05	395	2.62087E-05		-395	-2.87285E-06	395	2.87285E-06		-395	-9.99994E-05	395	9.99994E-05		-395	-5.0272E-06	395	5.0272E-06		-395	-9.11225E-06	395	9.11225E-06		-395	-9.99996E-05	395	9.99996E-05		-395	-3.14519E-06	395	3.14519E-06		-395	-4.11697E-06	395	4.11697E-06		-395	-4.20606E-05	395	4.20606E-05		-395	-1.36559E-05	395	1.36559E-05		-395	-1.00384E-05	395	1.00384E-05		-395	-1.83196E-05	395	1.83196E-05
-400	-2.70757E-05	400	2.70757E-05		-400	-3.04871E-06	400	3.04871E-06		-400	-9.99993E-05	400	9.99993E-05		-400	-5.24006E-06	400	5.24006E-06		-400	-9.67005E-06	400	9.67005E-06		-400	-9.99994E-05	400	9.99994E-05		-400	-3.1608E-06	400	3.1608E-06		-400	-4.17901E-06	400	4.17901E-06		-400	-5.65111E-05	400	5.65111E-05		-400	-1.46508E-05	400	1.46508E-05		-400	-1.04405E-05	400	1.04405E-05		-400	-1.89918E-05	400	1.89918E-05
-405	-2.78298E-05	405	2.78298E-05		-405	-3.2319E-06	405	3.2319E-06		-405	-9.99996E-05	405	9.99996E-05		-405	-5.40151E-06	405	5.40151E-06		-405	-1.02637E-05	405	1.02637E-05		-405	-9.99994E-05	405	9.99994E-05		-405	-3.17965E-06	405	3.17965E-06		-405	-4.36179E-06	405	4.36179E-06		-405	-7.52E-05	405	7.52E-05		-405	-1.56819E-05	405	1.56819E-05		-405	-1.08493E-05	405	1.08493E-05		-405	-1.95379E-05	405	1.95379E-05
-410	-2.84917E-05	410	2.84917E-05		-410	-3.42272E-06	410	3.42272E-06		-410	-9.99994E-05	410	9.99994E-05		-410	-5.63096E-06	410	5.63096E-06		-410	-1.09644E-05	410	1.09644E-05							-410	-3.19813E-06	410	3.19813E-06		-410	-4.44112E-06	410	4.44112E-06		-410	-9.85926E-05	410	9.85926E-05		-410	-1.67484E-05	410	1.67484E-05		-410	-1.12622E-05	410	1.12622E-05		-410	-2.0321E-05	410	2.0321E-05
-415	-2.95855E-05	415	2.95855E-05		-415	-3.61762E-06	415	3.61762E-06		-415	-9.99994E-05	415	9.99994E-05		-415	-5.82403E-06	415	5.82403E-06		-415	-1.17177E-05	415	1.17177E-05							-415	-3.21996E-06	415	3.21996E-06		-415	-4.45775E-06	415	4.45775E-06		-415	-9.99992E-05	415	9.99992E-05		-415	-1.7875E-05	415	1.7875E-05		-415	-1.16905E-05	415	1.16905E-05		-415	-2.09712E-05	415	2.09712E-05
-420	-3.02911E-05	420	3.02911E-05		-420	-3.82036E-06	420	3.82036E-06		-420	-9.99996E-05	420	9.99996E-05		-420	-6.22063E-06	420	6.22063E-06		-420	-1.26829E-05	420	1.26829E-05							-420	-3.23831E-06	420	3.23831E-06		-420	-4.55613E-06	420	4.55613E-06		-420	-9.99983E-05	420	9.99983E-05		-420	-1.9039E-05	420	1.9039E-05		-420	-1.21254E-05	420	1.21254E-05		-420	-2.17422E-05	420	2.17422E-05
-425	-3.12511E-05	425	3.12511E-05		-425	-4.02285E-06	425	4.02285E-06		-425	-9.99993E-05	425	9.99993E-05		-425	-6.65276E-06	425	6.65276E-06		-425	-1.40012E-05	425	1.40012E-05							-425	-3.25941E-06	425	3.25941E-06		-425	-4.76556E-06	425	4.76556E-06		-425	-9.99985E-05	425	9.99985E-05		-425	-2.02576E-05	425	2.02576E-05		-425	-1.25673E-05	425	1.25673E-05		-425	-2.23582E-05	425	2.23582E-05
-430	-3.22035E-05	430	3.22035E-05		-430	-4.23702E-06	430	4.23702E-06		-430	-9.99994E-05	430	9.99994E-05		-430	-7.01058E-06	430	7.01058E-06		-430	-1.55941E-05	430	1.55941E-05							-430	-3.27925E-06	430	3.27925E-06		-430	-4.86688E-06	430	4.86688E-06		-430	-9.99991E-05	430	9.99991E-05		-430	-2.15783E-05	430	2.15783E-05		-430	-1.30224E-05	430	1.30224E-05		-430	-2.29626E-05	430	2.29626E-05
-435	-3.30953E-05	435	3.30953E-05		-435	-4.47915E-06	435	4.47915E-06		-435	-9.99994E-05	435	9.99994E-05		-435	-7.6413E-06	435	7.6413E-06		-435	-1.75134E-05	435	1.75134E-05							-435	-3.29936E-06	435	3.29936E-06		-435	-4.93948E-06	435	4.93948E-06		-435	-9.99989E-05	435	9.99989E-05		-435	-2.29714E-05	435	2.29714E-05		-435	-1.34852E-05	435	1.34852E-05		-435	-2.37874E-05	435	2.37874E-05
-440	-3.39834E-05	440	3.39834E-05		-440	-4.77358E-06	440	4.77358E-06		-440	-9.99994E-05	440	9.99994E-05		-440	-8.69023E-06	440	8.69023E-06		-440	-1.95509E-05	440	1.95509E-05							-440	-3.31881E-06	440	3.31881E-06		-440	-5.143E-06	440	5.143E-06		-440	-9.99996E-05	440	9.99996E-05		-440	-2.44753E-05	440	2.44753E-05		-440	-1.39554E-05	440	1.39554E-05		-440	-2.43802E-05	440	2.43802E-05
-445	-3.50835E-05	445	3.50835E-05		-445	-5.51803E-06	445	5.51803E-06		-445	-9.99995E-05	445	9.99995E-05		-445	-9.15102E-06	445	9.15102E-06		-445	-2.19142E-05	445	2.19142E-05							-445	-3.33941E-06	445	3.33941E-06		-445	-5.29565E-06	445	5.29565E-06		-445	-9.99987E-05	445	9.99987E-05		-445	-2.60459E-05	445	2.60459E-05		-445	-1.44345E-05	445	1.44345E-05		-445	-2.52396E-05	445	2.52396E-05
-450	-3.59899E-05	450	3.59899E-05		-450	-7.2305E-06	450	7.2305E-06		-450	-9.99996E-05	450	9.99996E-05		-450	-1.08706E-05	450	1.08706E-05		-450	-2.46142E-05	450	2.46142E-05							-450	-3.35912E-06	450	3.35912E-06		-450	-5.4565E-06	450	5.4565E-06		-450	-9.99984E-05	450	9.99984E-05		-450	-2.7729E-05	450	2.7729E-05		-450	-1.49245E-05	450	1.49245E-05		-450	-2.60618E-05	450	2.60618E-05
-455	-3.69365E-05	455	3.69365E-05		-455	-9.74758E-06	455	9.74758E-06		-455	-9.99996E-05	455	9.99996E-05		-455	-1.17799E-05	455	1.17799E-05		-455	-2.74633E-05	455	2.74633E-05							-455	-3.38393E-06	455	3.38393E-06		-455	-5.70165E-06	455	5.70165E-06		-455	-9.99986E-05	455	9.99986E-05		-455	-2.95318E-05	455	2.95318E-05		-455	-1.54203E-05	455	1.54203E-05		-455	-2.66434E-05	455	2.66434E-05
-460	-3.78343E-05	460	3.78343E-05		-460	-1.24812E-05	460	1.24812E-05		-460	-9.99998E-05	460	9.99998E-05		-460	-1.37223E-05	460	1.37223E-05		-460	-3.0792E-05	460	3.0792E-05							-460	-3.40492E-06	460	3.40492E-06		-460	-5.90367E-06	460	5.90367E-06		-460	-9.99996E-05	460	9.99996E-05		-460	-3.15718E-05	460	3.15718E-05		-460	-1.59345E-05	460	1.59345E-05		-460	-2.7721E-05	460	2.7721E-05
-465	-3.86331E-05	465	3.86331E-05		-465	-1.5107E-05	465	1.5107E-05		-465	-9.99996E-05	465	9.99996E-05		-465	-1.51083E-05	465	1.51083E-05		-465	-3.45007E-05	465	3.45007E-05							-465	-3.43323E-06	465	3.43323E-06		-465	-6.01711E-06	465	6.01711E-06		-465	-9.99982E-05	465	9.99982E-05		-465	-3.41864E-05	465	3.41864E-05		-465	-1.64494E-05	465	1.64494E-05		-465	-2.81709E-05	465	2.81709E-05
-470	-3.95181E-05	470	3.95181E-05		-470	-1.76934E-05	470	1.76934E-05		-470	-9.99996E-05	470	9.99996E-05		-470	-1.85528E-05	470	1.85528E-05		-470	-3.79586E-05	470	3.79586E-05							-470	-3.45351E-06	470	3.45351E-06		-470	-6.22667E-06	470	6.22667E-06		-470	-9.99997E-05	470	9.99997E-05		-470	-3.77471E-05	470	3.77471E-05		-470	-1.69781E-05	470	1.69781E-05		-470	-2.93537E-05	470	2.93537E-05
-475	-4.05658E-05	475	4.05658E-05		-475	-2.01431E-05	475	2.01431E-05		-475	-9.99994E-05	475	9.99994E-05		-475	-1.83851E-05	475	1.83851E-05		-475	-4.17292E-05	475	4.17292E-05							-475	-3.48217E-06	475	3.48217E-06		-475	-6.41736E-06	475	6.41736E-06		-475	-9.9999E-05	475	9.9999E-05		-475	-4.13753E-05	475	4.13753E-05		-475	-1.75178E-05	475	1.75178E-05		-475	-3.00725E-05	475	3.00725E-05
-480	-4.15002E-05	480	4.15002E-05		-480	-2.23851E-05	480	2.23851E-05		-480	-9.99994E-05	480	9.99994E-05		-480	-2.26649E-05	480	2.26649E-05		-480	-4.5913E-05	480	4.5913E-05							-480	-3.5076E-06	480	3.5076E-06		-480	-6.85167E-06	480	6.85167E-06		-480	-9.99993E-05	480	9.99993E-05		-480	-4.50915E-05	480	4.50915E-05		-480	-1.80688E-05	480	1.80688E-05		-480	-3.10456E-05	480	3.10456E-05
-485	-4.24845E-05	485	4.24845E-05		-485	-2.43285E-05	485	2.43285E-05		-485	-9.99995E-05	485	9.99995E-05		-485	-2.7358E-05	485	2.7358E-05		-485	-5.00484E-05	485	5.00484E-05							-485	-3.53551E-06	485	3.53551E-06		-485	-8.38904E-06	485	8.38904E-06		-485	-9.99994E-05	485	9.99994E-05		-485	-4.87553E-05	485	4.87553E-05		-485	-1.86166E-05	485	1.86166E-05		-485	-3.19238E-05	485	3.19238E-05
-490	-4.36507E-05	490	4.36507E-05		-490	-2.60361E-05	490	2.60361E-05		-490	-9.99994E-05	490	9.99994E-05		-490	-3.0342E-05	490	3.0342E-05		-490	-5.44654E-05	490	5.44654E-05							-490	-3.56762E-06	490	3.56762E-06		-490	-1.12745E-05	490	1.12745E-05		-490	-9.99994E-05	490	9.99994E-05		-490	-5.26881E-05	490	5.26881E-05		-490	-1.91896E-05	490	1.91896E-05		-490	-3.25309E-05	490	3.25309E-05
-495	-4.52905E-05	495	4.52905E-05		-495	-2.76412E-05	495	2.76412E-05		-495	-9.99995E-05	495	9.99995E-05		-495	-3.66345E-05	495	3.66345E-05		-495	-5.92801E-05	495	5.92801E-05							-495	-3.59217E-06	495	3.59217E-06		-495	-1.53783E-05	495	1.53783E-05		-495	-9.99991E-05	495	9.99991E-05		-495	-5.70188E-05	495	5.70188E-05		-495	-1.97667E-05	495	1.97667E-05		-495	-3.35378E-05	495	3.35378E-05
-500	-4.70975E-05	500	4.70975E-05		-500	-2.92742E-05	500	2.92742E-05		-500	-9.99996E-05	500	9.99996E-05		-500	-4.57274E-05	500	4.57274E-05		-500	-6.47402E-05	500	6.47402E-05							-500	-3.61208E-06	500	3.61208E-06		-500	-2.30144E-05	500	2.30144E-05		-500	-9.99992E-05	500	9.99992E-05		-500	-6.25349E-05	500	6.25349E-05		-500	-2.03573E-05	500	2.03573E-05		-500	-3.43411E-05	500	3.43411E-05
20190227 Sensors (UCSC)
IV Curve
Bias (Voltage)
Leakage Current (A)