## Slide 1: ePixHR10k Low-Noise, High-Dynamic Range, High Rate Detectors for TXI
Electromechanical concept design status
April 2nd, 2021
Dionisio Doering*
*On behalf of SLAC teams

## Slide 2: April 2nd TXI camera meeting agenda
Review goals and requirements
Concept design evolution
  Mechanical / cooling
  Electronics
Teams
Schedule

## Slide 3: April 2nd TXI camera meeting agenda
Review goals and requirements
Concept design evolution
  Mechanical / cooling
  Electronics
Teams
Schedule

## Slide 4: Requirements Table
Parameter​ | Threshold​ | Objective​ | ePixHR 2M​
Pixel Pitch (um)​ | 100​ | 100​ | 100​
Read Noise (e- rms)​ | 110​ | 80​ | 75​
Frame Rate (kHz)​ | >5​ | 10​ | 5​
Array size (pixels)​ Forward Scattering Detector Von Hamos/Low-Q | >2MPix​ >110kPix | 4MPix​ 135kPix | 2.6MPix (2x1536x864)​ 110kPix (384x288)
Well Depth (Number of 4keV photons)​ | 10000​ | 10000​ | >16000​
Quantum Efficiency (% 2-4keV) | 80 | 90 | 95
Vacuum Compatibility | < 10-7 | < 10-7 | < 10-7
Cabling and Cooling Line Length | Compatible with LUSI Vacuum Detector Stage Model number: SA-391-520-00. | Compatible with LUSI Vacuum Detector Stage Model number: SA-391-520-00 | Compatible with LUSI Vacuum Detector Stage Model number: SA-391-520-00
Physical Package Envelope |  |  | 
Maximum Power Dissipation | <400W | <350W | 200W
Positional Stability | 0.1 pixel | 0.1 pixel | 0.1 pixel
Lowest Primary Vibration Mode | ‘Level D’ | ‘Level D’ | ‘Level D’
Y jaw movement | 40mm | 40mm | 40mm
Z translation | 0.5m | 0.5m | 0.5m

## Slide 5: Requirements Table
Parameter​ | Threshold​ | Objective​ | ePixHR 2M​
Pixel Pitch (um)​ | 100​ | 100​ | 100​
Read Noise (e- rms)​ | 110​ | 80​ | 75​
Frame Rate (kHz)​ | >5​ | 10​ | 5​
Array size (pixels)​ Forward Scattering Detector Von Hamos/Low-Q | >2MPix​ >110kPix | 4MPix​ 135kPix | 2.6MPix (2x1536x864)​ 110kPix (384x288)
Well Depth (Number of 4keV photons)​ | 10000​ | 10000​ | >16000​
Quantum Efficiency (% 2-4keV) | 80 | 90 | 95
Vacuum Compatibility | < 10-7 | < 10-7 | < 10-7
Cabling and Cooling Line Length | Compatible with LUSI Vacuum Detector Stage Model number: SA-391-520-00. | Compatible with LUSI Vacuum Detector Stage Model number: SA-391-520-00 | Compatible with LUSI Vacuum Detector Stage Model number: SA-391-520-00
Physical Package Envelope |  |  | 
Maximum Power Dissipation | <400W | <350W | 200W
Positional Stability | 0.1 pixel | 0.1 pixel | 0.1 pixel
Lowest Primary Vibration Mode | ‘Level D’ | ‘Level D’ | ‘Level D’
Y jaw movement | 40mm | 40mm | 40mm
Z translation | 0.5m | 0.5m | 0.5m
Primary requirements under investigation with the present concept design study

## Slide 6: Shingled 2Mpix ePixHR 5kHz camera
ePix10k / CSPAD envelope (OD. 286mm,
depth 175mm)
2 jaws with gap of
up to 40mm
Overall dimensions compatible with requirements
Each jaw composed of 2 abut quads
Each module composed of 4 sensor/ASICs (24 per detector)
20W - less than ePixM
Alternative solution can be a geometry with 5 tiles / quad
  Design for 6 populate 5.

## Slide 7: Drawing from experiences
Direct related detectors developments
  ePix10kHR single
    Each tile is equivalent to a single HR camera
      Power
      Data volume
  ePix 10k Analog detector
    Same pixel matrix
    Similar geometry (exception is the jaw)
  ePixM HR concept design
    Analog shingle design concept
    Less restrictive for the ePix10k
      Higher temperature
      More space
TXI detector
ePix10k HR single
ePixM HR shingle concept
ePix10k Analog detector

## Slide 8: April 2nd TXI camera meeting agenda
Review goals and requirements
Concept design evolution
  Mechanical / cooling
  Electronics
Teams
Schedule

## Slide 9: Proposed concept
4 independent quadrants to build 1 camera
  Power
  Data
2 independent halves
  Mechanics bonds these 4 quads
11 distinct boards
  Tile
  Flex (6 different lengths)
  Interposer (left/right) board
  Digital (left/right) board

## Slide 10: Tile specs
Each module composed of 4 sensor/ASICs (24 per detector)
1 Sensor
4 ASICs
Ceramic substrate for ASIC/Sensors cooling
Tiles are using a set of stringers
  Mechanical support
  Cooling path
PCB
144*192 pixels * 4 ASICs = 110kpix

## Slide 11: PPT to CAD
Alternative solution can be
  Geometry with 5 tiles / quad
  Reduce max gap
  Increase camera envelope
2.6MPix (2x1536x864)​

## Slide 12: Reducing the number of tile to 5
Alternative solution can be
  Geometry with 5 tiles / quad
  Reduce max gap
  Increase camera envelope
2.2MPix (2x1536x720)​
* Does not include overlap of tiles

## Slide 13: Concept design with 6 tiles – Multiple views
Front view of the detector
Back side of the tiles showing single connector per time
Front side of the tiles, showing shingle concept

## Slide 14: Tile step size
Optimizations are under way
  Based on simulation of the cooling for the tiles
  Option selection will be based on thermal analyses
    Ceramic substrate thickness
    Followed by step size
  Shingle overlap distance
    At the moment, it is set to zero
    Will be increase to minimize gaps
    Depends on the gap size
6.35mm/5mm
4.57mm/2mm
4.57mm/1.5mm

## Slide 15: April 2nd TXI camera meeting agenda
Review goals and requirements
Concept design evolution
  Mechanical / cooling
  Electronics
Teams
Schedule

## Slide 16: Tile to Interposer/analog board
Baseline until now was
  6 tile rows
  Evaluating impact on 5 tile rows
Connectors
  Tile connector to interposer
  Interposer to digital board
  Power connector
Tile support electronics
  LDO, ADC, DAC, temp. sensor
20mm
143mm
768
864
1536
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active area
Balcony
144
192
123mm

## Slide 17: Tile to Interposer/analog board
Baseline until now was
  6 tile rows
  Evaluating impact on 5 tile rows
Connectors
  Tile connector to interposer
  Interposer to digital board
  Power connector
Tile support electronics
  LDO, ADC, DAC, temp. sensor
20mm
143mm
768
720
1536
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active area
Balcony
144
192
123mm
* Envelope violation will go away when overlap is defined.

## Slide 18: ePix HR Interposer quad board
Interposer left and right quads
Connectors
  Tile connector to interposer
  Interposer to digital board
  Power connector
Tile support electronics
  LDO, ADC, DAC, temp. sensor
20mm
123mm
143mm
Power
Data flow

## Slide 19: ePix HR Digital Quad board
Digital left and right quads
Single board per quad
  One large FPGA
  FO
  Local power
Functionalities
  ASIC control, configuration and data readout
  Support electronics control
  Data communication to DAQ
143mm
Power?
FO
20mm
123mm

## Slide 20: Power estimation
Power supply voltage level
  2.5V @ 2.1A
LVDS
  12 IOs for data transmission per ASIC
  2 IOs clock receiver
single ended
  17 control lines inputs and output
  1 x Analog monitor
Power supply voltage level
  2.5V @ 8.4A
LVDS
  72 data transmission per module
  3 clock receiver
single ended
  20 control lines inputs and output
  3 x Analog monitor
Power supply voltage level
  2.5V @ 50A
LVDS
  432 data transmission per camera
  6 clock receiver
single ended
  68 control lines inputs and output
  18 Analog monitor
1 ePix 10k ASIC(5.2W)
4 ePix 10k ASIC(20.8W)
48/40 ePix 10k ASIC (249.6W/216W) / ½ camera
or
24/20 ePix 10k ASIC (124.8W/108W) / ¼ camera

## Slide 21: Data flow
FPGA
+
HBM*
50Gbps
Mon.
ADC
Power supply
MM - FO
OSCs
LCLS-II timing (Optional)
Control & Mon. link (1Gbps)
(Optional)
SM - FO
…
x6 ASICs carriers and concentrator cards
*High Bandwidth Memory can be used to improve on camera processing. Example VUP37 with 8GB HBM
Each ASICs carrier board is connected to a concentrator card
Concentrator I/O requirements:
  48 IO/module => 288 IO/Quad
  (6 x 2 IOs per line x 4 ASICs x 1 modules)
  44 (Ideal) / 23 (w/o monitors) IO/module for ASIC control =>130 IO/Quad
  12 Analog monitoring ADC IO /module
  8 Digital monitoring IO/module
Data volume at 5kFPS
  8.5Gbps/module
  (144*192 pixels * 4 ASICs * 16 b * 5kFPS, 66b/64b)
  51Gbps/quad (6 modules per quad)
  42.5Gbps/quad (6 modules per quad)
Data transmission 50Gbps/40Gbps* out
  2x25Gbps lanes/4x10Gbps lanes
  Assume a lossless compression will be used.
ASIC
ASIC
ASIC
ASIC
ASIC
ASIC
ASIC
ASIC
ASIC
ASIC
ASIC
ASIC

## Slide 22: April 2nd TXI camera meeting agenda
Review goals and requirements
Concept design evolution
  Mechanical / cooling
  Electronics
Teams
Schedule

## Slide 23: Working groups
Overall mechanics
  Andy, Serge, MMcK and design team (1/month)
  external interface, requirements, constraints
  Jaw movement
Internal electro-mechanical system
  Chris K., Angelo D., Martin N., Marco O., Maciej K., Bojan M., Umanath K., Steve B. and Dionisio D.) weekly
  Mechanical support and cooling
  Tile, analog and digital board design
Software/Firmware weekly (not started yet)
  Maciej K., Larry R., Dionisio D.
  Data collection, aggregation, pre-processing and transmission

## Slide 24: April 2nd TXI camera meeting agenda
Review goals and requirements
Concept design evolution
  Mechanical / cooling
  Electronics
Teams
Schedule

## Slide 25: Milestones and Deliverables
Key Milestones | Planned Date | Goal
Conceptual Design Review | 01/13/2021 - Today | Project High level review
Kick-off meeting | 02/19/2021 | Start design
Mechanical Design Review | 05/15/2021 | Structural and cooling design review
Electro-mechanical design concept | 05/31/2021 | Detector concept development including hutch integration
Preliminary Design Review | 06/01/2021 | Concept review
Detailed Electro-mechanical design completed | 10/20/2021 | Complete Detector design
Final Design Review | 10/21/2021 | Manufacturability review
Components fabrication starts | 10/22/2021 | Parts fabrication and procurement
Firmware Design completed | 12/08/2021 | Fully operational boards
Components fabrication completed with qualification | 01/07/2022 | Parts fabricated and qualified
Modules and Detector Assembled | 08/22/2022 | Assembled deliverables
Detector Characterization and Integration | 08/23/2022 | Parameter optimization and hutch integration
Detector Delivery | 09/21/2022 | Transfer to operations
Project closeout | 09/22/2022 | 
Deliverables:
1 ePixHR10k 2.6MPix (or 2.1MPix) with a 2 jaws design at 5KHz
2 ePixHR10k 110kPix at 5kHz (single planar module)
New!

## Slide 27: Electromechanical interface
https://confluence.slac.stanford.edu/display/ppareg/AIR+X-ray+Detectors
Resources page about the project

## Slide 28: Discussion topics
Progress report on mechanical design concept

## Slide 29: ePix HR quad Tile – Board view
40mm
70mm
Shortest (mated) flex 40mm with 10mm flexible length.
Longest (mated) flex 70mm with 40 mm flexible length.
SEAM8-50-S02.0-S-04-2
SEAF8-50-1-S-04-2-RA
Connectors:
50 mm
Active Area
Active Area
Active Area
Active Area
46 mm
15.5 mm
15.5 mm
77mm

## Slide 30: ePix HR quad Tile – Ceramic substrate view
40mm
70mm
Shortest (mated) flex 40mm with 10mm flexible length.
Longest (mated) flex 70mm with 40 mm flexible length.
SEAM8-50-S02.0-S-04-2
SEAF8-50-1-S-04-2-RA
Connectors:
50 mm
Active Area
Active Area
Active Area
Active Area
46 mm
15.5 mm
15.5 mm
77mm

## Slide 31: Space available
40mm
70mm
Shortest (mated) flex 40mm with 10mm flexible length.
Longest (mated) flex 70mm with 40 mm flexible length.
SEAM8-50-S02.0-S-04-2
SEAF8-50-1-S-04-2-RA
Connectors:

## Slide 32: 1 Schematic, 2 layout (left and right)
143mm
Power
20mm
123mm
143mm
Data flow
Data flow
Power

## Slide 33: Tile to Interposer/analog board
Connectors
  Tile connector to interposer
  Interposer to digital board
  Power connector
Tile support electronics
  LDO, ADC, DAC, temp. sensor
20mm
123mm
143mm
768
864
1536
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active Area
Active Area
Active Area
Active Area
Active Area
Active Area
Balcony
Active area
Balcony
144
192
85mm
77mm
46mm
50mm
5mm

## Slide 34: ePix HR Interposer quad board
Interposer left and right quads
Connectors
  Tile connector to interposer
  Interposer to digital board
  Power connector
Tile support electronics
  LDO, ADC, DAC, temp. sensor
20mm
123mm
143mm
Power
Data flow

## Slide 35: ePix HR Digital Quad board
Digital left and right quads
Single board per quad
  One large FPGA
  FO
  Local power
Functionalities
  ASIC control, configuration and data readout
  Support electronics control
  Data communication to DAQ
20mm
123mm
143mm
77mm
85mm
46mm
5mm
50mm
Power?
FO

## Slide 36: ASIC characteristics
Power supply voltage level
  2.5V @ 2.1A
LVDS
  12 IOs for data transmission per ASIC
  2 IOs clock receiver
single ended
  17 control lines inputs and output
  1 x Analog monitor
Power supply voltage level
  2.5V @ 8.4A
LVDS
  72 data transmission per module
  3 clock receiver
single ended
  20 control lines inputs and output
  3 x Analog monitor
Power supply voltage level
  2.5V @ 50A
LVDS
  432 data transmission per camera
  6 clock receiver
single ended
  68 control lines inputs and output
  18 Analog monitor
3 ASICs(21W)
18 ASICs (126W)
1 ePix 10k ASIC(5.2W)
4 ePix 10k ASIC(20.8W)
48 ePix 10k ASIC(249.6W) / ½ camera
or
24 ePix 10k ASIC(124.8W) / ¼ camera
1 ASIC(7W)
ePixM

## Slide 37: Infrastructure required from support electronics
ASIC
  Data connectivity
  Power regulated at 2.5V
  Control (SACI + Control lines such as clock, resets, start of readout, etc.)
DACs
  Bias DACs
  1 DAC for external (ramp) ADC tests
ADCs
  Fast ADCs
    Multiplex input, fast for the analog monitors (?)
    Can be multiplexed, or can skip some ASICs
  Slow ADCs for power and current consumptions (must have)
Transceivers
  High speed communication (50gpbs/quad minimum, consider 250gpbs)
Memory
  Internal HBM (desirable feature)